## survey
type | name | label::en | required | required_message::en | required_message::fr | appearance | relevant | default | constraint_removed | constraint_message::en | choice_filter | calculation | hint::en | media::image::en | media::video | instance::type | instance::tag
begin group | inputs | Patient |  |  |  | field-list | ./source = 'user' | [9 empty] | hidden
string | source | Source |  |  |  | hidden |  | user
string | source_id | Source ID |  |  |  | hidden
begin group | user | User
end group
begin group | contact | Contact
string | name | Name |  |  |  | hidden
string | date_of_birth | Date of birth |  |  |  | hidden
string | sex | Sex |  |  |  | hidden
integer | external_id | Muso ID
end group
end group
calculate | patient_id | Patient ID | yes | This field is required | Ce champ est obligatoire | [6 empty] | ../inputs/contact/_id |  |  |  |  | hidden
calculate | has_symptoms | Has Symptoms | yes | [8 empty] | if(${symptoms},’true’,’false’) |  |  |  |  | hidden
calculate | test_result | Test Result | [9 empty] | if(selected(${interpretation},’c’),if(selected(${interpretation},’m’) or selected(${interpretation},’g’),’positive’,’negative’),’invalid’) |  |  |  |  | hidden
begin group | symptoms_group | Patient Assessment for symptomatic contacts |  |  |  | field-list | [10 empty] | hidden
select_multiple select_symptoms | symptoms | Is your patient experiencing any of the following symptoms?
note |  | <style>.pages .loader { display: none; }</style>
end group
begin group | precautions_group | Precautions |  |  |  | field-list | ${has_symptoms}=’true’ | [9 empty] | hidden
select_multiple select_precaution | precautions | Please confirm that you have taken the following precautions: | yes | [5 empty] | selected(., ‘ppe’) and selected(., ‘disinfect’) and selected(., ‘splashes’) | Please confirm each of the safety precautions before proceeding
end group
begin group | preparations_group | Preparation |  |  |  | field-list | ${has_symptoms}=’true’ | [9 empty] | hidden
select_multiple select_preparation | preparations | Check and prepare the test kit | yes | [5 empty] | selected(., ‘expiry’) and selected(., ‘open’) | Please confirm each of the preparation steps before proceeding
end group
begin group | finger_group | Executing the test |  |  |  | field-list | ${has_symptoms}=’true’ | [9 empty] | hidden
select_multiple select_finger | finger | Prick the patient’s finger | yes | [5 empty] | selected(., ‘disinfect’) and selected(., ‘prick’) and selected(., ‘flow’) | Please confirm each of the test steps before proceeding
note | [13 empty] | finger.png
end group
begin group | disinfect_group | Executing the test |  |  |  |  | ${has_symptoms}=’true’
note |  | Collect blood into capillary tube or sampler (~20-50uL) | [11 empty] | collect_blood.png
note |  | Put the blood into the sample well of the cassette | [11 empty] | blood_to_well.png
note |  | After the blood has completely entered the sample well, add 2 to 3 drops (70 to 100μl) of buffer solution. Make sure to test on a level surface at room temperature. | [11 empty] | buffer_to_well.png
end group
begin group | wait_group | Set timer for 15 minutes |  |  |  | field-list | ${has_symptoms}=’true’ | [9 empty] | hidden
note | evaluation_timer | Begin 15 minute timer | yes |  |  | countdown-timer |  | 900
select_multiple select_wait | waiting | While you’re waiting: | [6 empty] | selected(., ‘wash_hands’) and selected(., ‘discard’) | Please confirm these steps before proceeding
end group
begin group | interpret_group | Interpret COVID-19 Test Result |  |  |  | field-list | ${has_symptoms}=’true’ | [9 empty] | hidden
note | [13 empty] | interpretation_guide.png
select_multiple select_interpret | interpretation | Which lines appear on your test?
end group
begin group | cleanup_group | Test Cleanup |  |  |  | field-list | ${has_symptoms}=’true’ | [9 empty] | hidden
select_multiple select_cleanup | cleanup | Now that you’ve completed the COVID-19 test, please cleanup your work area | yes | [5 empty] | selected(., ‘dispose’) and selected(., ‘clean’) | Please confirm you have completed each cleanup step before proceeding
end group
begin group | summary-group | Summary of Assessment Result |  |  |  | field-list
note | no_symptoms | This assessment is for patients exhibiting symptoms of COVID-19. Your patient did not have symptoms of COVID-19 and should not be tested. <ul><li>Please advise your patient to stay home and limit contact with others.</li><li>Your patient should call or text if symptoms develop</li><li>Follow-up with your patient in one week</li></ul> |  |  |  |  | ${has_symptoms}=’false’
note | positive_result | This assessment concluded <b>your patient is positive for COVID-19</b>, please take these <b>immediate actions</b>: <ul><li>Accompany the patient to the health facility for clinical test and care.</li><li>Keep a safe physical distance from the patient and wear a mask and gloves to protect yourself.</li><li>Help the patient keep a safe physical distance (greater than 2 meters) from others while traveling.</li></ul> |  |  |  |  | ${test_result}=’positive’
note | negative_result | <b>The test result for your patient is negative.</b> But this doesn’t mean they do not have COVID-19. Since hey were showing symptoms, you should advise the following: <ul><li>Advise your patient to stay home and avoid seeing other people if possible.</li><li>Your patient should call or text if symptoms worsen or new symptoms arise.</li><li>Follow up with your patient in 24 hours.</li></ul> |  |  |  |  | ${test_result}=’negative’
note | invalid_result | <b>No conclusion can be determined about your patient</b>. The RDT test that you used for this assessment was invalid. Please submit this form, but restart a new COVID-19 Assessment with a new RDT test. |  |  |  |  | ${test_result}=’invalid’ and ${has_symptoms}=’true’
note | style | <style>[name='/covid_rdt/summary-group'] ul { margin: 10px 0 0 25px; list-style: disc; }</style>
end group

## choices
list_name | name | label::en
select_precaution | ppe | Put on protective equipment. This should include a gown, gloves, and surgical mask.
select_precaution | disinfect | Disinfect your work surface before starting work
select_precaution | splashes | If possible, place an alcohol soaked towel on your work surface to catch splashes
select_preparation | expiry | Select a test kit and check the expiration date. Tests should not be used past their expiration date.
select_preparation | open | Open the test pouch and prepare to do the test. After opening, tests must be used within one hour.
select_interpret | c | C
select_interpret | m | M
select_interpret | g | G
select_cleanup | dispose | Dispose the cassettes, pipettes as biohazard materials.
select_cleanup | clean | Clean work surfaces, and plastic materials (pipettors, pens, and timer) with alcohol at the end of the work.
select_wait | discard | Discard capillary sampler into the trash
select_wait | wash_hands | Wash your hands with 70% alcohol (keep gloves on)
select_wait | talk | Have a conversation with your patient and their family about regular hand washing, symptom monitoring, and social distancing.
select_symptoms | fever | Fever
select_symptoms | cough | New cough
select_symptoms | breathing | Difficulty breathing
select_symptoms | diarrhea | Diarrhea
select_finger | disinfect | Disinfect the patient’s finger, use their middle or ring finger
select_finger | prick | Prick the patient's finger to begin the flow of blood
select_finger | flow | Apply light pressure, wipe away the first 2-3 drops of blood with alcohol pad, and make sure there is free blood flow

## settings
form_title | form_id | version | style | path | instance_name
COVID-19 Assessment | covid_rdt | 2020-03-22 18-38 | pages | data